IGSTK: Image-Guided Surgical ToolKit

An Open Source platform for Facilitating Image-Guided Surgical Application Development

Kevin Cleary

1 Dec 2005

# Motivation

The Image-Guided Software Toolkit (IGSTK: pronounced IGstick) is an open-source project aimed at supporting the development of reliable software for patient-critical medical applications, particularly image-guided surgery applications.

Image-guided surgery involves the use of images to provide instrument guidance and to aid in clinical decision making during procedures. Using intra-operative computation allows pre-operative images to augment or eliminate the need for intra-operative images during image-guided procedures. Such computer-based image-guided systems have been commercially available for almost ten years now, but developing new image-guided systems is an active field of research in the medical imaging community. Since these systems are intended to be used for patient care in a hospital setting, the reliability of the software is paramount. The majority of the effort in developing a new image-guided system is debugging and validating the software for implementing the algorithms, controlling the system, and displaying the results. The process we have evolved in creating this toolkit allows a distributed group of developers to collaborate with a focus on testing and quality control.

These considerations led us to form a multidisciplinary team of software engineers, medical imaging scientists, and clinical personnel to develop the image-guided surgery toolkit IGSTK. Requirements were developed through interviews with Interventional Radiologists who specialize in image-guided procedures. Three example applications were selected to drive the requirements process. The toolkit contains the basic software components of an image-guided system along with the example applications and a framework for application development. Using this toolkit and our software process we believe that it is possible to create reliable and extensible software in an open source project that helps address FDA requirements and that can be used by researchers in academia and commercial companies.

# Image-Guided Surgery

## Background

Image-guided surgery is a rapidly developing field, since these procedures typically mean substantially less trauma for the patient. Computer-based image guidance was originally developed for neurosurgical applications as navigation in the brain requires great precision (Watanabe, 1987). This technology allows the physician to use pre-operative computed tomography (CT) or magnetic resonance imaging (MRI) scans to guide minimally invasive procedures. Image guidance was then extended to spinal and orthopedic applications, such as pedicle screw placement (Foley and Smith 1996). Commercial image-guided surgery systems are now available for brain, spine, and ENT applications.

The basic concept of image-guided surgery is to provide the physician with a real-time update of the anatomy located in the region of a surgical instrument. Typically, this capability is provided by tracking the location of the instrument and the patient using devices known as trackers. By also designating the locaton of patient trackers in the pre-operative images it is possible to display the location of the instrument on the pre-operative image. This display can help guide the physician to the precise anatomical target and also provides the physician with a type of “x-ray vision” in that the physician can see what lies beneath a surgical instrument before starting the procedure.

A typical image-guided system is a complex merger of three major components: 1) a control computer; 2) software for image processing, control, and the user interface; and 3) a tracker for localizing instruments and/or the patient in three-dimensional space. Figure 1 shows a typical commercial computer-aided surgery system, the Stealthstation Treon™ from Medtronic Sofamor Danek.  On the left hand side is the optical localizer, and on the right hand side is the control computer and display.

Figure 1: Typical image-guided surgery system

(courtesy of Medtronic Surgical Navigation Technologies)

The following is a typical sequence of steps in using an image-guided surgery system:

- A pre-operative CT or MRI scan is obtained – fiducials may be placed on the anatomy before this scan for later use in registration during the procedure.

- The CT or MRI images are imported into the computer – DICOM is the most common format used, although many vendors support proprietary formats as well.

- A “reference target” is attached to the anatomy to compensate for any inadvertent motion of the camera or patient.

- Registration is the next step – this procedure maps the image data set to the patient’s location and anatomy.

- The system can now track surgical instruments, including probes or pointers, and display the anatomy beneath these instruments.  A typical four quadrant view (axial, saggital, coronal, and 3D) is shown in Figure 2.

- Multiplanar reconstructions (i.e. oblique reformats) can also be provided at any angle to help carry out the procedure.

Figure 2: Typical four quadrant display (axial, saggital, coronal, and 3D)

(courtesy of Medtronic Surgical Navigation Technologies)

## Need for Robust Software

The critical component of this system which consumes the great majority of the development effort when creating a new image-guided system is the software. The software must integrate information from multiple trackers, correlate this information with image information, and display real-time updates of the instruments and anatomy. The components are very tightly coupled and can be difficult to debug and validate using functional programming style.

Image-guided and minimally invasive techniques are continuing to grow in popularity and the international research community is rapidly expanding as well. However, because there is no readily available open-source software for image-guided surgery, many research groups are forced to develop their own software at a substantial cost both in time and effort. The main contribution of this project to the research community is the delivery of a high quality software toolkit (IGSTK) suitable for in the development and validation of new image-guided surgery applications.

## Example Applications from Toolkit

{note that here I just pasted the text from the Phase II proposal – obviously this will need to be greatly edited to maybe just one page and I will do so after I discuss with Kevin Gary and others}

### Task 1: Demonstration Applications

Once the toolkit has been implemented and tested in Aim 2, the next step will be to develop three demonstration applications. Each of the three collaborating organizations (Kitware, Georgetown, and UNC) will be primarily responsible for one of these applications. The goal of this aim is to show the usefulness of the toolkit in developing clinically related applications as well as determining if there are any shortcomings in the toolkit that need to be rectified. These applications are intended to test the suitability of the IGSTK toolkit for implementing real clinical applications. These applications are not the final goal of this project and parts of these applications are being developed under independent funding so we are leveraging that effort here. The applications are:

- Application 1: Guidewire tracking (Georgetown)

- Application 2: Needle segmentation and tracking (Kitware)

- Application 3: Liver biopsy (UNC)

#### Application 1: Guidewire Tracking (Georgetown)

The goal of this application is to assess the feasibility of device navigation with electromagnetic tracking through a phantom study. We will provide an image overlay of an electromagnetically tracked guidewire as it is moved through the vasculature. The user interface will be similar to the example application shown in Section C.3.1.4, but we will also provide a magnified “scout or scanogram” view of the phantom to show the general position of the guidewire in the phantom, along with an axial CT scan of the phantom corresponding to the tip of the guidewire. This would be useful in clinical applications such as aneurysm clipping, stent graft deployment and positioning, and catheter navigation for chemoembolization or transhepatic portal systemic shunt creation, as it would provide the physician with a more detailed view of the anatomy.

The clinical partner for this application will be Brad Wood, MD, an interventional radiologist at the NIH Clinical Center. Dr. Wood has collaborated with the principal investigator Dr. Cleary on a number of other projects and an abstract on the guidewire tracking concept has been accepted at the 2004 Society of Interventional Radiology annual meeting. This application will be validated on the vascular phantom developed by Dr. Wood in conjunction with Georgetown and shown in Figure 3. The phantom includes several vessels that can accept the magnetically tracked guidewire. The guidewires will be supplied by Neil Glossop, PhD, of Traxtal Technologies as part of another project (see letter of support in the appendix). A picture of a prototype guidewire is shown in Figure 4.

| Figure 3: Vascular phantom for guidewire test | Figure 4: Electromagnetically tracked guidewire (courtesy of Traxtal Technologies) |
| --- | --- |

Once the software has been developed, the utility of the system will be evaluated as follows. This test will be done in the interventional suite at Georgetown. The sequence of steps is:

- Fiducials that are visible in CT images are placed on the phantom. Internal fiducials along the vascular tree (such as major bifurcations and identifiable branches or plaques) will also be identified and used in a separate arm of the phantom study.

- The phantom is CT scanned.

- The images are loaded on the workstation and the fiducials are then located in CT space by asking the interventionalist to select them using the mouse.

- The fiducials are located in electromagnetic space by enabling the electromagnetic tracking system and touching each fiducials with the magnetically tracked guidewire.

- The information from steps 3 and 4 is used to determine the transformation between CT space and electromagnetic space using the method of Arun (Arun, Huang et al. 1987).

- The guidewire can then be tracked on a scout image to show the general position of the guidewire within the phantom as well as showing an axial CT image at the tip of the guidewire.

- With the phantom in the angiography environment, the guidewire will be placed in a catheter in the vessels in the phantom, and the utility and error of the system will be evaluated by two interventional radiologists (Dr. Wood and Dr. Levy). Navigation will be performed using angiography with dual display of tracking position in CT space (axial and multiplanar reconstruction images based upon position and orientation of guidewire).

#### Application 2: Needle Segmentation and Tracking (Kitware)

Needle detection and tracking is a task commonly found in image-guided surgery. A set of IGSTK components will be implemented for providing the functionalities of needle segmentation and tracking. These components will be exercised in an application used for guiding the positioning of a surgical needle based on the information provided by an intra-operative fluoroscopy system.  The needle tracking application considered here has been already developed at the Georgetown ISIS center under independent funding with the aim of guiding a robotically assisted needle insertion system. This application will be reworked in order to factorize the software elements that will constitute needle tracking component. The application will be adapted in order to use the new component and verify that it performs as intended.

Figure 5 illustrates the clinical environment in which this application is currently being tested. A fluoroscopy C-Arm is used for gathering intra-operative images of the patient. A robot is used for positioning the tip of the surgical needle at the entry point in the patient’s skin and for orienting the needle in order to aim at a target in the internal anatomy.  In this procedure, the final insertion of the needle is performed by the clinician by using a joystick to control the robot. Figure 6 shows the segmentation of the surgical needle overlaid into the fluoroscopic image. The blue points indicate the pixels segmented as members of the needle. The cross marks are references to the region of interest used during the segmentation and to the position of the needle’s head, tip and middle points respectively.

| Figure 5: Clinical setup for robotically assisted needle insertion. | Figure 6: Surgical needle segmentation from a fluoroscopic image. | Figure 7: Fluoroscopic image of needle aligned with the direction of insertion. |
| --- | --- | --- |

In order to calibrate the position of the robot with respect to the image-plane of the fluoroscope, we displace the robot arm along known directions and by known distances in a 2D plane parallel to the operating table. These displacements, when projected onto the image plane of the fluoroscope result in a different set of directions and distances as seen on the fluoroscopic image. By measuring the observed displacements in the image plane it is possible to estimate a two dimensional affine transform mapping the robot coordinate system into the fluoroscopic image coordinate system.  This mapping disregards the perspective effect due to the natural divergence of X-rays emanated from the fluoroscope’s source. This approximation is justified due to the small region in space under which the robot is displaced, typically a cube with side length of 50mm. Once the 2D affine transformation between the robot and image planes has been estimated, it is then used for computing the movement required in the robot arm in order to reach a target that has been identified in the fluoroscopic image. Because of the lack of depth information in a single projection of the fluoroscopic image, the needle tip is simply placed on the line of sight between the anatomical target and the fluoroscope. Having the needle’s tip on the line of sight, the application proceeds to orient the needle until it is parallel to the line of sight. The needle is considered to be aligned when the size of its projection on the fluoroscopic image is minimized. This geometrical configuration is shown in Figure 7. At that moment the needle is ready for insertion in the patient’s anatomy. The C-Arm is then rotated to a position where the line of sight is almost orthogonal to the direction of insertion, and the control is left to the clinician for manually control the insertion process using the robot’s joystick. This application has been validated in a preliminary cadaver study were BBs (small metal balls) have been inserted to serve as verifiable targets.

Figure 8 illustrates the IGSTK software components that can be integrated in this application in order to provide high level functionalities. It shows a video input component used for acquiring the fluoroscopic image, a needle segmentation component for detecting the needle on the image, and a 3D/2D registration component intended to overlay geometrical models of the anatomy in a virtual view of the fluoroscopic image.

Figure 8: IGSTK Components used to modularize the implementation 
of the robotically assisted needle placement application.

#### Application 3: 3D Ultrasound Augmentation for Percutaneous Liver Lesion Biopsy Guidance (UNC)

This project demonstrates the extensibility of IGSTK to guide liver lesion radiofrequency ablation (RFA) procedures by fusing intra-operative ultrasound images with pre-operative CT and MR images.  The complete application, built using IGSTK, compensates for organ movement without requiring intravensous access and is generally useful when organ movement cannot be well predicted or tracked using external markers.

In liver lesions readiofrequency ablation, the lesion is localized and then an RFA probe (needle) is advanced through the skin and into the lesion to the desired location. Once the needle is positioned, it is attached to a generator and current is sent through the needle and into the tumor according to the manufacturer’s protocol. Temperatures of 100 C lasting 10-30 minutes are obtained to insure cell death.  At UNC, we use Berchtold RFA devices that deliver burns as large as 7 cm that allow the treatment of a 5 cm spherical lesion with a 1 cm “negative margin” burn around the lesion. Because of the desired 1 cm safety margin around each lesion, our radiologists desire to position an RFA needle within 5 mm of their intended location. While needle placement accuracy is not solely determined by the accuracy of RFA site localization from the pre- to the intra-operative images, we nevertheless maintain 5 mm as the maximum acceptable registration error for our system.  Other groups have stated similar requirements [King 2001a,b].

Two general IGSTK modules are being developed for this application.  First, a 3D ultrasound image composition module constructs 3D ultrasound volumes from tracked 2D ultrasound probes.  Second, an image-based registration module performs image-to-image registration to align the intra-operative 3D ultrasound data with the pre-operative MR/CT data.

The 3D ultrasound image composition module is an IGSTK image source.  It involves 3D tracking of a 2D ultrasound probe.  The 2D images are then painted into a 3D volume.  This module addresses the challenging task of reconstructing a valid liver image despite liver movement due to respiration, cardiac motion, and probe pressure during the acquisition of a sequence of 2D ultrasound images.

The image-based registration module accounts for the internal movement of organs. This module

builds upon the AURORA module to track the ultrasound probe and patient in order to initialize the registration of the intra-operative 3D ultrasound data and the pre-operative CT/MR data.  The image-based registration module then refines that registration so as to account for the intra-operative displacement of a designated organ.  This modules builds upon the image-based registration framework in ITK.  The user can substitute different transforms (e.g., rigid, affine, or deformable), different image registration quality metrics (e.g., mutual information or cross-correlation), and different optimization algorithms (e.g., conjugate gradient or Levenberg-Marquardt).

The tracked needle is then overlaid onto the fused ultrasound/CT/MR data using existing IGSTK modules.  Figure 9 illustrates the fusion of a lesion seen on CT with vessels and data from an intra-operative 3D ultrasound scan – the fused view is oriented with respect to the position of the patient on the table.   Figure 10 illustrates preliminary results from the registration metric: one slice of co-registered MR and 3D ultrasound data of a liven (Aylward, Bullitt et al. 1996; Bullitt, Liu et al. 1999; Aylward and Bullitt 200), ITK, and image registration (Aylward, Jomier et al. 2002; Aylward, Jomier et al. 200) experience.  Our particular expertise is the development of clinically useful software.

| Figure 9: 3D ultrasound data is oriented as the patient is oriented on the table to simplify biopsy guidance. Vessels, tumor, and RFA probe are shown as 3D surface renderings. | Figure 10: One slice of fused 3D US and MR liver data. The liver vessels and surfaces, visible in both images, are well aligned despite respiratory movement. |
| --- | --- |

(end of text)

[10] E. Watanabe, T. Watanabe, S. Manaka, Y. Mayanagi, and Y. Takakura, “Three dimensional digitizer neuronavigator: New equipment for computed tomography-guided stereotactic surgery,” Surg. Neurol., vol. 27, pp. 543–547, 1987.

Also known as computer-aided surgery in the field.

DICOM stands for Digital Imaging and Communications in Medicine and has become the standard medical image file format 5.	DICOM reference guide. Health Devices, 2001. 30(1-2): p. 5-30..  All major medical vendors of CT and MRI scanners provide a DICOM capability.