IGSTK: Framework and Example Applications Using an Open Source Toolkit for Image-Guided Surgery Applications

Kevin Clearya*, Peng Chenga, Hui Zhanga, Hee-su Kimb, Kevin Garyc, M. Brian Blaked, David Gobbie, Stephen Aylwardf, Julien Jomierf, Andinet Enquobahrief, Rick Avilaf, Luis Ibanezf

aImaging Science and Information Systems (ISIS) Center, Dept. of Radiology, 
Georgetown University, Washington, DC, USA

bKyungpook National University, Korea

cDivision of Computing Studies, Arizona State University, Mesa, Arizona, 85212, USA

dDepartment of Computer Science, Georgetown University, Washington, DC, 20007, USA

eAtamai Inc., London, Ontario, N6B 2R4, Canada

fKitware Inc., Clifton Park, NY, 12065, USA

ABSTRACT

Open source software has tremendous potential for improving the productivity of research labs and enabling the development of new medical applications. The Image-Guided Surgery Toolkit (IGSTK) is an open source software toolkit based on ITK, VTK, and FLTK, and supports common operating systems such as Linux, Windows, and MacOS. It integrates basic components needed in surgical guidance applications and provides a common platform for fast prototyping and development of robust image-guided applications. This paper gives an overview of the IGSTK framework and current status of development followed by an example needle biopsy application to demonstrate how to develop a surgical application using this toolkit.

Keywords: open source, surgical guidance, tracking, state machine, needle biopsy, application prototyping

1. PURPOSE

The field of image-guided surgery is rapidly expanding, as image-guided surgical techniques are being developed that minimize treatment invasiveness and thereby reduce trauma and speed recovery for patients. A typical image-guided system consists of tracking subsystems for capturing the movement of instruments and the patient and a computer subsystem for integrating and displaying pre- and intra-operative images and instrument positions using the tracking data. The software of the computer subsystem usually requires the majority of the development effort. As new trackers and data integration algorithms are developed, each research group must develop its own image-guided system, which can be a time-consuming undertaking. Testing such systems can also be an extensive process, with risks to patients and surgeons. The purpose of IGSTK is to enable the creation and validation of reusable, robust image-guided software as open-source, free of commercial licensing restrictions. This component based software toolkit intends to make it easy for researchers to prototype and develop image-guided surgery applications.

# 2. METHODS

IGSTK has been under development for over a year and a public beta release will be available by the time of the conference, February 2006. IGSTK is based on the following existing opening source software toolkits: ITK for segmentation and registration [1], VTK for visualization [2], and FLTK for the user interface [3]. The toolkit contains basic software components needed to construct an image-guided system, including a support for trackers, data integration algorithms, and a four-quadrant view incorporating image overlay.

Given that IGSTK is intended to be used for developing applications that will run in the surgical room, robustness and quality are the highest priorities in the design of the toolkit. To minimize the risk of harm to the patient resulting from misuse of the classes, IGSTK is designed based on the following principles [4]:

- Requirements are generated by studying the surgical procedure in clinical environment. Only features that are necessary for providing functionalities requested in the requirements are implemented. This is essential for ensuring the quality and safety of the toolkit as it is intend to be used in critical applications.

- All IGSTK components are encapsulated with state machine. State machine is contained within the class to control the access to the class. Components are always in a valid state, so it will perform in a predicable manner. The use of state machine will also enforce high quality standards for code coverage and run-time validation.

- IGSTK does not expose ITK and VTK objects on APIs and unwanted feature in those objects are being rewritten and encapsulated. This is to prevent user get access to the feasibility of those objects.

- IGSTK does not appreciate feasibility. Instead it creates explicit objects and performs type checking through API. In that way, a minimum number of parameters are left free for the application developers to define, and therefore the opportunities for error are drastically reduced, and also prevent the miss use of objects.

- IGSTK use event to communicate between components, event evoked by one state machine could be mapped into input to another state machine.

- IGSTK does not use return value or exception at application level. Instead it uses event to pass information around.

## 2.1 State machine

To ensure the safety and robustness, state machine design pattern is incorporated into the IGSTK components. A state machine is defined by a set of states, a set of inputs, and a set of directed transitions between states. Transitions are changes from one state to another when certain stimulus or input is received. An action may be taken along with a transition, to perform some task on entry to a state, or on exit from a state. State machines support the component-based architecture through encapsulation and extensibility. It ensures that a component is always in a valid state which can guarantee repeatable and deterministic behavior. Detailed description on the state machine theory can be found in the literature [6].

The following points summarize the advantage of using state machine in the toolkit [7]: {need rework on this paragraph, it’s just a copy and paste from IEEE paper}

- Safety and Reliability: A state machine ensures that component behavior is deterministic and that all components are in a known and error-free state at any given moment.

- Cleaner design: Since developers must anticipate all possible inputs, states, and transitions, the state machine encourages and enforces a cleaner and more robust design, free of untested assumptions.

- API simplicity: A focused, clearly expressed application programming interface (API) is a must for supporting robustness and reliability. In the context of surgical guidance, we believe that flexibility and abundance of features are undesirable, because they create more opportunities for things to go wrong during a surgical intervention.

- A consistent integration pattern:  The toolkit’s value as it matures will undoubtedly be tied to the incorporation of additional functionality at the component level.  This functionality will often take the form of reusable code from existing toolkits.  State machines provide a consistent pattern for integrating this functionality while adhering to the safety-first principles necessitated by the application domain.

- Quality Control: State machines facilitate code coverage in the sense of lines of code tested, as well as path coverage on a per component basis. Using code that is not based on state machines may result in applications that exhibit unreliable behavior.  At run time, they can easily enter into any number of untold states that were never explored by the developer, and can lead to error conditions that may or may not become visible to the users of the application.

The safety emphasis in IGSTK encourages explicit knowledge of component state and determines whether a given behavior may be executed while in that state.  Hence, our application is a combination of traditional concepts in reliability (being in a known state) and the State Pattern [8] (for managing availability of behavior dynamically).

2.2 Architecture and components

Figure 1 presents a UML collaboration diagram of the major IGSTK components involved in a typical image-guided surgery application. The key components are from left to right [4, 5]:

- View classes are used for presenting results and displays to the physician. An internal pulse generator is used to update the scene at specified refresh rate to keep the visual information up-to-date. Currently the IGSTK View classes are built with VTK classes and FLTK window. They take FL events and translate them into VTK events in order to enable user interaction within the render window. View2D and View3D are subclasses of View with different interaction style.

- SpatialObject and SpatialObjectRepresentation.  SpatialObjects classes model physical objects including image data and simple geometrical shapes such as anatomical structures and surgical devices, while SpatialObjectRepresentations classes are responsible for displaying its associate SpatialObject in the scene. IGSTK Spatial Objects can be attached to a Tracker Tool object, so that it can update its position as the Tracker Tool moves. The state machine in SpatialObjects classes enforces this association to be one time operation, once a Spatial Object is attached to a tracker tool it is not expected to get back to manual nor to be re-attached to a second tracker tool. It is no longer possible to change the position of the object programmatically. IGSTK has a variety of SpatialObjects and corresponding SpatialObjectsRepresentations, including basic shapes (Cone, Cylinder, Ellipsoid and so on), ImageSpatialObject, UltrasoundProbeObject, and GroupObject to make composite SpatialObjects.

- Tracker classes are abstraction of tracking devices. They communicate with hardware and store data including position, orientation, and other relevant information pertaining to tracked surgical tools. The IGTSK Tracker component is developed based on code donated by Atamai Inc. and warped with state machine. The internals of the Tracker require a separate thread to communicate with tracking hardware at an acceptable frequency. An pulse generator inside the class is used to automatically update the location information of the tracked device with expiration time and propagate this changes to SpatialObjects. IGSTK now supports NDI AURORA, POLARIS, and VICRA trackers (Northern Digital Inc., Waterloo, Canada) and Flock of Birds tracker (Ascension Technology Corporation, Burlington, VT USA).

Figure 1: IGSTK Component Architecture

Some other components do not present in Figure 1 include: DICOMImageReader for read in DICOM images, Landmark3DRegistration for 3D pair-point-based rigid body registration, Logger for logging debug or other information, and Calibration class for calibrating tracker tools. All these together provide essential functionalities to build an image-guided system.

# 3. Example Application

This section gives an example application developed for the needle biopsy procedure. For a typical image-guided needle biopsy procedure, a pre-operative CT image is acquired and later registered to the patient using landmark fuducials. During the biopsy phase, the needle is being tracked by some optical tracking device with real-time visualization of its location overlaid on top of the CT image providing guidance to the surgeon for better targeting the needle to the desired location. In this example, state machine is being implemented at the application level. However this is not mandatory, using state machine gives the application developer an easier way to prototype the application and to control the work flow of the surgical procedure, and also adds an extra layer of security to the application to prevent it to perform unwanted tasks when it’s not in the right state. In the following sections, we are going to show you how to write an application under the IGSTK frame work.

3.1 Prototyping

By analyzing a typical needle biopsy procedure, we identify the following essential tasks in this procedure [9].

- Register the patient information.

- Load in the pre-operative CT image. Fiducials (small markers) are usually placed on the anatomy prior to the scan for landmark based registration in Step 4~7.

- Verify the registered patient information against that loaded from image. Prompt the surgeon if there is a discrepancy. This step is performed to prevent operation on the wrong site.

- Identify the image landmarks by going through the CT image slices and click on the fuducials in the image. For pair-point based registration, at least three points are required.

- Initialize the tracking device.

- Add patient landmarks by touching the physical fuducials attached to patient using the tracked pointer device.

- Perform the image to patient landmark registration.

- Path planning. The surgeon will select a target point and an entry point to plan the path for the needle puncture.

- Real-time display of the overlay of the needle probe and pre-operative images in the quadrant viewer window during the biopsy procedure.

Thinking of the whole application as a state machine, the completion of each task will cause the application enter a new state, so we have a set of states to indicate current status of the application. The user interaction on the GUI can be translated into inputs to the state machine. For instance, when we click on the register patient button, this will generate a “RequestSetPatientNameInput” to the state machine. The state machine will take this input and change its current state from “InitialState” to “WaitingForPatientNameState”, and the action is to pop up a window asking for input of the patient name. If the user input a valid name, then there will be a “PatientNameInput” which brings the state machine into “PatientNameReadyState”, otherwise there will be a “PatientNameEmptyInput”, which will return the state machine to the “InitialState”. Now we can map the application into series of states and inputs, this higher level abstraction will help the developer design a clear work flow for the application. Figure 2 shows the state machine diagram for the needle biopsy application generated automatically when state machine is constructed. This intuitive graphic illustration of the state machine will facilitate the developer with the prototyping process and work flow design. Some recurring patterns on state machine design can be found on http://public.kitware.com/IGSTKWIKI/index.php/State_Machines#Recurring_Patterns_in_State_Machine_Design.

Figure 2: State machine diagram for needle biopsy application. (Black for states and blue for inputs)

3.2 Coding the state machine

In this section we are going to show how to code the state machine into the needle biopsy application. IGSTK has a number of convenient macros to facilitate the programming of the state machine. The details of these macros can be found in Source/igstkMacro.h. More information on state machine design pattern and guidelines could be found on http://public.kitware.com/IGSTKWIKI/index.php/State_Machines and http://public.kitware.com/IGSTKWIKI/index.php/State_Machine_Design_Guidelines.

Once we have the higher level abstraction of the application and prototyped it in the state machine model, we need to take the following steps to program the state machine into the applications.

1. The first thing to do is use the state machine declaration macro in your class’s head file. What this macro does is to define types for state and input, creates a private member variable m_StateMachine, and two private member functions for exporting the state machine description into dot format for the state machine diagram visualization [10] and LTSA (Labeled Transition Systems Analyzer) format for state machine animation and validation [11].

igstkStateMachineMacro();

2. Then take the states and inputs we mapped out during the prototyping stage, and define them in the head file using the following macros. To enforcing the name conventions of the state machine, the declaration macros will append “State” or “Input” automatically after the variable name. For instance, the following two lines will define m_InitialState and m_RequestLoadImageInput.

igstkDeclareStateMacro( Initial );

igstkDeclareInputMacro( RequestLoadImage );

3. Next step is to construct the state machine in the source file, in the constructor. First, we need to add all the states and inputs declared in the header into the state machine.

igstkAddStateMacro( Initial );

igstkAddInputMacro( RequestLoadImage );

4. The next step is the crucial step, we are going to setup the state machine transition table, which essentially controls the whole logic and work flow of the application. This is being done by using the following macro, igstkAddTransitionMacro( From_State, Received_Input, To_State, Action). This means when the state machine is in the “From_State” and receive the “Received_Input”, it will enter into the “To_State” and evoke the ActionProcessing() as an action for this transition. This macro requires the “Action” method to be pre-defined in the class for the state machine to call. For example:

igstkAddTransitionMacro( Initial, RequestSetPatientName, WaitingForPatientName, SetPatientName );

In this case, we need to have a SetPatientNameProcessing() method defined in the class to have this code compiled.

5. After we have setup the transition table, the next step is to select an initial state, and flag the state machine to be ready to run. After the state machine is being set ready to run, we can not change the state machine transition table in the code, this is being designed in a way to enhance the safety of the state machine and prevent accidentally change the state machine behavior manually in the code.

igstkSetInitialStateMacro( Initial );

m_StateMachine.SetReadyToRun();

6. Now the state machine is being setup and ready to run. We can then export the state machine description in dot format and generate the graphical visualization as shown in Figure 2. This graph will help us to examine the work flow of the application and the state transition table.

std::ofstream ofile;

ofile.open("DemoApplicationStateMachineDiagram.dot");

const bool skipLoops = false;

this->ExportStateMachineDescription( ofile, skipLoops );

ofile.close();

This will output the state machine into a dot file when we execute the application. If you have dot tool installed in your system, then you can run following command, which will take the dot file and generate png format picture named “SMDiagram.png” for the state machine.

>dot –T png –o SMDiagram.png DemoApplicationStateMachineDiagram.dot

7. All the requests to a state machine should be translated into inputs and the state machine will response to those inputs depending on its current state. These actual actions should be protected method and only be called by state machine directly. So in the code, a click on the load image button will be translated in to a ‘RequestLoadImageInput’, and then we call ProcessInputs(), let the state machine to handle this request.

igstkPushInputMacro( RequestLoadImage );

m_StateMachine.ProcessInputs();

If the state machine is in the right state to load the image, a protected method associated with this transition (eg. LoadImageProccessing() ) will be evoked by the state machine as defined in the transition table constructed in the constructor.

3.3 Discussion

From the computational theory point of view, all computers are state machines, and all computer programs are state machines regardless of whether the developers used the state machine programming pattern or not. Traditional programming approaches represent the states ambiguously by using a large number of variables and flags, which results in lots of conditional test in the code (if-else or switchcase statements in C/C++). Programmers could easily forget to consider all possible paths in the code while struggling with the if-else conditional test and flags check. The result of those approaches is unpredictable behavior and unsafe designed applications, which is definitely unwanted for mission critical applications running in the surgery room.

Compare to the traditional approaches, state machine will radically reduce the number of path in the code, save the developers from convoluted conditional test, and encourage them focusing on higher level design. From the above example, we can see the state machine is easy to program and manage under the IGSTK frame work. We encourage developers design and code the state machine of their application first, and then generate the state machine diagram as shown in Figure 2. They can go through the figure, examine and verify their design of the work flow. If they want to add or change a path of the application, it is just a matter of adding or deleting a transition table entry. This is much easier than going through the code and struggle with those if-else logics. This will largely facilitate the application prototyping, and the implementation code could be plugged into the skeleton program later. All these will result in clearer designs and safer applications.

Figure 3 shows the user interface of the needle biopsy application written in FLTK. The left side is the control panel, consists of a set of buttons corresponding to the series of tasks performed during the procedure. These buttons’ callbacks should call the public request methods of the application, which will be translated into state machine inputs. And the state machine will take proper reaction according to its own state. For example, when the patient information is not set, the ‘Load Image’ button won’t respond to the user click. There is no need for conditional check or disabling of buttons here, they are already in the state machine transition table. On the right hand side, there are four standardized views, axial, sagittal, coronal, and 3D view. Here we loaded an abdominal phantom CT image. The ??? is the planned path, and the green cylinder represents the needle being tracked by the tracker. The viewer will automatically reslice the images as the needle tip moving in the anatomy {Update this part when new image is available}.

Figure 3: Snapshot of the needle biopsy application {Will use another image with tracked objects}

# 4. CONCLUSIONS

The major components of the toolkit have been completed. Automated nightly testing results can be found at http://public.kitware.com/dashboard.php?name=igstk. The current code coverage on tests is around 95.7%. A state machine validation tool is being developed using LTSA [11], which can animate the transition of the state machine. This tool can be found on the wiki page http://public.kitware.com/IGSTKWIKI/index.php/IGSTK_State_Machine_Validation.

Further development of this tool is to test all possible combination of inputs to a state machine, which will allow automated testing of applications implemented with state machine. We are planning on releasing three demo applications with this toolkit. The “Needle Biopsy” application as illustrated in the previous sections is delivered. “Ultrasound Guided RFA” and “Guidewire Tracking” will be available before the stable release in June 2006.

The IGSTK toolkit is open source software distributed under a BSD-like license. Basic information related to IGSTK can be found at the website www.igstk.org and the Wiki pages http://public.kitware.com/IGSTKWIKI/. Instructions for configuring and building the toolkit are available at http://public.kitware.com/IGSTKWIKI/index.php/How_to_build_IGSTK. You are welcome to try the software, review the source code and send your comments and report bugs to the IGSTK user list, igstk-users@public.kitware.com. Note that this software should only be use in clinical cases under IRB approval. You are allowed to use IGSTK for free in academic and commercial applications but it is your responsibility to perform the tests and validations required by regulatory bodies such as the U.S. Food and Drug Administration (FDA).

ACKNOWLEDGMENTS

This research is supported by the National Institute of Biomedical Imaging and Bioengineering (NIBIB) at the National Institutes of Health (NIH) under grant R42EB000374 and by U.S. Army grant W81XWH-04-1-0078. The content of this manuscript does not necessarily reflect the position or policy of the U.S. Government. The authors would like to thank the IGSTK advisory board for their advice throughout the project: Will Schroeder of Kitware; Ivo Wolf of the University of Heidelberg; Peter Kazanzides and Anton DeGuet from John Hopkins University; and Ingmar Bitter, Matt McAuliffee, and Terry Yoo from the NIH.

REfereNCES

- L. Ibanez and W. Schroeder, “The ITK Software Guide”, ISBN 1-930934-10-6, 2005 http://www.itk.org/ItkSoftwareGuide.pdf

- W. Schroeder and K. Martin, B. Lorensen, “The Visualization Toolkit, An Object Oriented Approach to 3D Graphics”, Kitware Inc., 1998

- FLTK a cross-platform C++ GUI Toolkit, http://www.fltk.org/documentation.php/doc-1.1/toc.html

- L. Ibanez, J. Jomier, D. Gobbi, R. Avila, M. B. Blake, H. Kim, K. Gary, S. Aylward, and K. Cleary. "IGSTK: A State Machine Architecture for an Open Source Software Toolkit for Image-Guided Surgery Applications." Insight journal, Aug 2005

- K. Cleary, L. Ibanez, S. Ranjan, and M.B. Blake. "IGSTK: A Software Toolkit for Image-Guided Surgery Applications." CARS 2004

- J.E. Hopcroft, et al., Introduction to Automata Theory, Languages, and Computation, Addison Wesley (2nd ed.), 2000.

- K. Gary, M. B. Blake, L. Ibanez, D. Gobbi, S. Aylward, and K. Cleary.” IGSTK: An Open Source Software Platform for Image-Guided Surgery.” Submit to IEEE Computer

- E.Gamma, et al., Design Patterns: Elements of Reusable Object-Oriented Software, Addison-Wesley, 1995.

- F. Banovac, N. D. Glossop, D. Lindisch, D. Tanaka, E. Levy, and K. Cleary: Liver Tumor Biopsy in a Respiring Phantom with the Assistance of a Novel Electromagnetic Navigation Device. MICCAI (1) 2002: 200-207

- Graphviz - Graph Visualization Software, http://www.graphviz.org

- LTSA - Labelled Transition System Analyser, http://www.doc.ic.ac.uk/~jnm/book/ltsa-v2

* cleary@georgetown.edu; Telephone 1-202-687-8253; Fax 1-202-784-3479