The image-guided surgery toolkit IGSTK: 
an open source C++ software toolkit

ABSTRACT

INTRODCUTION

1.1	Image Guided Medical Procedures - Background

Minimally invasive procedures are becoming increasingly popular in today’s healthcare system. Patients prefer these procedures to open surgeries since they cause fewer traumas to the body and faster recovery time.  Interventional radiologists and surgeons are also getting more experienced and becoming more comfortable performing these procedures.  The main driving forces behind the wide acceptance of minimally invasive procedures are advancements in medical imaging, registration algorithms, visualization technologies and tracking systems.

Medical imaging is a technique of creating image map of the human body.  Clinicians use medical images to diagnose diseases and perform actual treatments. Researchers also utilize medical imaging techniques to study normal human anatomy and physiological functions. Since the discovery of X-ray, medical imaging has been advancing at alarming rate. State of the art imaging techniques generate high quality and high resolution images. X-ray Fluoroscopy, Computed Tomography (CT), Magnetic resonance imaging (MRI), and ultrasound are the principal modalities used for minimally invasive procedures.   X-ray Fluoroscopy uses x-rays to produce 2D cross-sectional image of the body. In Computed Tomography (CT), reconstruction algorithms and computer technology are used to generate a 3D volume from 2D cross-sectional images generated. X-ray Fluoroscopy and CT are the preferred modalities for imaging bony structures. On the other hand, MRI is the principal imaging modality to image the brain due to its superb ability to discriminate between subtle differences in soft tissues (INSERT REFERENCES). MRI produces pictures of the internal body structures by detecting small changes in the frequency and phase of the radio-frequency signals emitted by resonant hydrogen nuclei of the tissues.  Recently, ultrasounds are commonly being used in minimally invasive procedures.  Ultrasounds use high frequency sound waves and their echoes to produce cross sectional images of the body. Through the years image quality of ultrasound has advanced to make it useful for diagnosis and surgery guidance.

To use medical images for surgical navigation, we need to first correspond them with the patient’s anatomy. This step is called registration. It is only when we can accurately register the image coordinate system with the patient coordinate system that we can take the full advantages of high quality images. Registration is the technique of computing spatial transform that maps points from one coordinate system to another.  Registration has many applications in image guided surgery procedures. These include registration of pre-operative image to intra-operative image, registration pre-operative images to patient coordinate system and registration of images from different modalities. For example, X-ray Fluoroscopy which accounts for more than 90% of intra-operative image for guidance generates 2D cross-sectional images. Interventional radiologist has to mentally register these 2D cross-sectional images to 3D patient body to perform the procedure. This can create ambiguity and inaccuracies in the procedure. However, by registering a pre-operative 3D MRI or CT image to the X-ray fluoroscopic image, a rich 3D image information can be generated and used in the procedure. This will result in a better performance.  Throughout the years, several registration algorithms have been developed. The primary approaches are image intensity based, landmark based (fiducial landmarks or anatomical landmarks), surface based and stereotactic frame based techniques. (INSERT REFERENCES)

The other integral component of image guided procedure is the tracking system. Tracking system tracks the real-time position of surgical tools and patient anatomy. Visual feedback can be provided by superimposing these changing positions of the surgeon’s tools on the preoperative or intra operative images.  This helps the surgeon navigate inside patient body by referencing to the overlaid images. Currently, there are three different tracker systems available. They are based on ultrasound, electromagnetic field and optical systems [INSERT REFERENCES]. In ultrasonic and optical instruments, there should be unobstructed line of sight between sensors of the instrument and transmitting device. Electromagnetic tracking systems don’t have this limitation but it is affected by ferromagnetic objects in its working volume. In recent years, accuracy of these tracking systems has been steadily improved contributing to the increased in the success rate of image guided procedures.

INSERT PICTURES OF OPTICAL AND MAGENETIC TRACKERS

To provide a visual feedback to the clinician, registered information of the tracking instrument and pre-operative or intra-operative image is presented to the clinician in a meaningful manner. Visualization plays a key role in this regard. The visualization application displays a real time virtual anatomy around where the instrument is located. Traditionally, slice-by-slice or multi-planar display view is used. Typically the images are reformatted to create a view perpendicular to the direction or plane of the probe. With increase in the resolution of medical images, surface and volume rendering techniques are commonly used to display 3D information of the surgical scene. Surface rendering techniques display surface structures identified within the image (the surface could be generated using contouring techniques).  Related rendering technique is texture mapping in which the original data is pasted on selective surfaces. For volume rendering, a ray-tracing technique is commonly used where the intensity of the image at each point is a function of the characteristics of the structures traversed by the ray.  Advancement in these algorithms and computing computer has made interactive rendering of 3D scenes possible in surgical procedures.

INSERT A PICTURE OF 3D RENDERED IMAGE WITH SURGICAL INSTRUMENT OR TRACKING PROBE (Visualization technology demonstration)

Image guided medical procedures are broadly classified into two groups: Image guided intervention and image guided surgery. With increasing overlap in the techniques and clinicians cross-performing some of the procedures, the distinction is becoming increasingly blurred.

In image guided intervention, the radiologists perform diagnostic and therapeutic procedures guided by images without surgery involved. Interventional radiologists mainly use instruments such as needles and catheters. Examples of diagnostic procedures are lesion biopsy and injection of contrast agents for CT angiograms imaging.  Therapeutic procedures include using stents to open clogged arteries and RF ablation techniques to destroy tumor tissues.

In image guided surgery procedures, surgeons perform surgery by inserting instruments in small incisions guided by pre-operative images. To track the position of the surgical instruments, a navigation system is used. The navigation system shows the position and orientation of the surgical system in the context of pre-operative images.  The preoperative image is aligned with the patient coordinate system using registration. The tracking system is used to track the position of the surgical instrument.  Although, the majority of image guided surgeries are performed on brain, application on other organs is not uncommon. For example, interventional MRI is frequently used to identify breast tumors and guide surgical excision of the tumor.

1.2	Use of Open source software in medical applications

The use of open source software in medical application is gaining more acceptances in recent years.  With the US government agencies such as NIH and NCI encouraging open source software development, more research groups are developing and releasing software as open source [INSERT REFERENCES]. Open source software permits users to study, use or modify the source code without any restrictions.

The Visualization Toolkit (VTK) and Insight Toolkit (ITK) are prime examples of the success of open source software projects. VTK [REFERENCE] originally developed at GE Global research provides a wide range of algorithms in computer graphics, image processing, and visualization. The toolkit has a wide user base with thousands of researchers and developers actively using it all over the world.  VTK provides advanced multi-dimensional visualization algorithms and modeling techniques such as implicit modeling, polygon reduction, mesh smoothing, cutting, contouring, and Delaunay triangulation and parallel computing.

ITK [REFERENCE] developed under the support of the US National Library of Medicine (NLM) at the National Institute of Health provides advanced registration and segmentation algorithms.  The toolkit contains the state of the art algorithms and it has been continually being developed with supports from NLM. Similar to VTK, the software is implemented in C++ programming language and provides wrapper for Tcl, Python and Java interpreters.

The composition of the ITK development team reflects the strength of a collaborative effort that is greatly emphasized in open source project. The initial development team consists of three commercial partners (GE Corporate R&D, Kitware Inc. and MathSoft) and three academic partners (University of Tennessee, University of North Carolina and University of Pennsylvania).  Other groups are also actively contributing to the toolkit.

Several benefits have contributed for increased popularity and acceptance of open source software. Some of the main ones are discussed below.

- Use of open source software encourages collaboration between medical researchers and commercial partners.

- Use of open source software avoids reinventing the wheel. Often times a researcher wastes valuable time and meager resources implementing basic infrastructures. With the availability of open source implementations of basic infrastructure, researchers will be able concentrate on the main scientific problems.

- Open source software provides valuable resource for educational purposes. The best way to learn software development is by studying how other developers have done it.

- Open source software provides implementation of reference or benchmark algorithms for validation and verification purposes.

- Open source software permits users to try out the technology without any cost restrictions and this will ultimately encourage rapid dissemination of the technology. Subsequently, this will tremendously accelerate research and development.

1.3	IGSTK project

IGSTK the focus of this paper is another open source software developed with a support from the National Institute of Biomedical Imaging and Bioengineering at NIH.  It is a cross-platform C++ library that provides basic components necessary to develop an image-guided system.  It is a powerful toolkit for fast prototyping and development of robust image guided applications.  The initial version of the software was released to the public in February 2006 at the SPIE Medical Imaging Conference in San Diego.

The toolkit was developed with collaboration between academic and industry partners.  Dr Kevin Cleary at the Imaging Science and Information Systems (ISIS) Center in Georgetown University is the principal investigator of the project. ISIS is a medical research and development institution that specializes in information and imaging technology in healthcare. The commercial partner is Kitware, Inc., a small company specializing in open-source software. Kitware also provides commercial consulting services for ITK and VTK. Expertise in tracking systems was provided by Atamai, Inc. (Ontario, Canada). Experts from Arizona University provide guidance in software process management and object-oriented technologies.

The reminder of the paper consists of four parts. Section II describes the software process adopted for IGSTK development. This will be followed by a discussion about the architecture.  Section IV outlines strategies being followed to build IGSTK user community. Finally, a short conclusion section is presented.

2	IGSTK Software process

IGSK is an open source project designed to develop mission critical application. This has necessitated the adoption of a software process that emphasizes on safety and robustness and at the same time light-weighted so that developers who are geographically separated and with various area and level of expertise can contribute to the toolkit efficiently. The adopted software process is very similar to agile unified process (INSERT REFERENCE) where light weight, iterative, incremental, test driven development approach is utilized.  IGSTK software process has four major components: version control system, build and release management, automated testing and communication and documentation.

Version Control System

Managing software version and update and is essential in iterative software development process. IGSTK uses Concurrent Versioning system (CVS). CVS follows a client-server model where a central server is maintained for developers to check out, modify and merge changes back to the code base. CVS allows simultaneous edits and branching and makes releases easier. IGSTK has two code repositories: main CVS and sandbox.  The main CVS repository contains version of the code that has been reviewed, well tested and approved for integration. The code in the main CVS adheres to all the quality software policies established in IGSTK. On the other hand, the sandbox is a testbed for developers to experiment on new ideas before integration with the rest of the toolkit and it is subjected to lesser quality requirements.

Build and Release Management system

Configuring and managing build process of a project is a challenging task in software development.  Large projects employ various tools and libraries. Developers use different compilers and development environments for coding. The build process should be able to coordinate all these tools. The task is particularly formidable for a cross-platform toolkit like IGSTK. Cross-platform project configuration task includes selecting appropriate compiler, finding required packages and libraries, specifying include header file and library paths and selecting the appropriate compiler flags and options consistently across platforms.

IGSTK build process is controlled by CMake [INSERT REFERENCE]. CMake is a cross-platform build system. CMake simplifies the configuration process by using platform independent configuration files to generate native build files such as UNIX makefiles and Visual Studio workspace targeting the user selected compiler. CMake automated configuration process makes it possible to develop a cross platform toolkit without too much additional effort.

IGSTK has public and internal release cycles. The public releases are scheduled twice a year. Public releases contain completed major functionalities of the toolkit. In between public releases, the team holds internal releases approximately two or three months.   Internal releases comprise implementation of new features on the components. Release information is posted in the IGSTK Wiki and the main website (www.igstk.org).

Automated Testing

Incremental development has to guarantee a continuous and stable software code base. Testing is critical to maintain the quality of software. Incremental development benefits from automated testing. Automated testing provides instant feedback of the impact of new additions to the code base or bug fixes to the rest of the toolkit.  IGSTK being a cross-platform software is tested and results reported to the developers on various combinations of compilers and platforms. This is made possible using the Dart system [INSERT REFERENCE]. Dart is an open-source toolkit designed for a distributed software quality testing. In the dart system, the software is run on distributed multiple combinations of compiler and platforms and the result is made available in publicly accessible a web page for review and feedback. In addition to compilation and build results, code coverage, unit test and dynamic analysis results are posted in the dashboard. This mechanism ensures a continued quality of the software even during development. The dashboard accepts nightly, experimental and continuous build results. Every night at a specified time, the entire toolkit is built and all the tests are run and results are posted in the dashboard.  During the day, if new changes are committed to the repository, continuous builds are submitted automatically to the dashboard. In this way changes are immediately tested without waiting for the nightly build and the developer can easily trace the problem if the new changes cause problems to the rest of the toolkit.  Experimental builds are used to test locally modified version of the toolkit. This provides the developer with a useful feedback about the impact of the local change to the rest of the toolkit before committing it to the repository.

INSERT A PICTURE OF A DASHBOARD

Communication and documentation

Communication is essential for the success of a collaborative work. IGSTK software process facilities direct and smooth communication among developers and project leads.

Weekly teleconference, idea sharing using Wiki and mailing list are found to be powerful tools of communication. Technical topics are discussed in great detail regularly during the tcons. Furthermore, the dashboard is reviewed at each tcon to maintain the quality of the software.

Documentation is essential for a continued use of a toolkit. IGSTK follows literate programming philosophy. Literate programming encourages developers to include a human-readable documentation into the source code file. IGSTK uses Doxygen tool [INSERT REFERENC] for automated document generation from the source code. Doxygen employs simple mark up language that is embedded in the source code. This mechanism has two main benefits. First, contrary to traditional software development approach, developers will not waste valuable time after code development to write documentations. The documentation will be ready when code development is finished. Secondly, the documentation will stay current with the code. Revision of the documentation after every software release won’t be necessary. In IGSTK, the document generation process is also integrated with the nightly build system. Every night the documentation is generated and made accessible from the dashboard. Hence, users have access to the latest documentation.

To summarize the software process, here are the ten best practices adopted by the IGSTK team [INSERT REFERENCE]

- Recognize that people are the most important mechanism available for ensuring high quality software:

- Facilitate constant communication among developers:

- Produce iterative releases:

- Manage source code from a quality perspective:

- Focus on 100% code and path coverage at the component level:

- Emphasize continuous builds and testing

- Support the development process with robust tools:

- Manage requirements iteratively in lockstep with code management:

- Focus on meeting exactly the current set of requirements:

- Evolve the development process:

3. IGSTK Architecture

IGSTK supports development of patient-critical image guided surgery applications. Hence the guiding principle in IGSTK architecture design is safety. Ensuring the safety of patients is the most critical aspect of the design. To achieve this safety-by-design, the IGSTK team followed component based architecture [INSERT REFERENCES] and state machine [INSERT REFERENCES] software design methodologies.   Each IGSTK component has a well defined set of features governed by state machines. The state machine ensures that the component is always on a valid state and all its transitions are valid and meaningful.

In component based architecture, the main functionalities are broken down to functional components with well defined interfaces for communication between the components. This type of architecture has several benefits in developing a reliable and robust toolkit. Breaking down the complex functionalities into smaller pieces makes implementation manageable. Consequently complexity and implementation details are encapsulated in the component level and hidden from the application developer or user. This encapsulation allows IGSTK to manage complexity inherent in image guided surgery. Furthermore, testing is easier at component level compared to traditional procedural software thereby increasing the quality of the software. Application developers can also use the components independently in their application without the need to integrate the full toolkit.  This will increase the adoption of the software and subsequently benefit the user community. Lastly, component based architecture is suitable for a structured toolkit extension. This will encourage users to extend the toolkit and contribute back to the community.

A state machine is defined by a set of states, set of inputs and set of directed transitions. Each IGSTK component is subjected to state machine control. Incorporation of state machines has tremendously enhanced the reliability of the toolkit for the following reasons. State machines ensure that the components have deterministic behavior at all times and are always in a known and error-free state. Formal validation of the software can be easily done since inputs, states and transitions are well defined and discrete.  This framework is suitable for automated testing. Interaction between the user and other applications can be explicitly defined using state machines reducing the error in application development. In summary, state machines ensure safety and reliability, cleaner design, API simplicity, consistent integration pattern and for quality control. State machines prevent wrong use of the state machine, mange complexity, traceability, and suitability for testing.

IGSTK components are implemented using C++ programming language. The component implementations are mostly based on ITK and VTK classes subjected to a strict state machine control.

The main components in IGSTK are the following:

• View (Display)

• Spatial Objects (Geometric Representation)

• Spatial Object Representation (Visual Representation)

• Tracking

• Readers

Short descriptions of these components are given below:

View (Display)

Viewers display the graphical representations of the renderings of the surgical scenes. Surgeons will be able to perform their task by visualization the information provided in the viewers.  Furthermore, the view components serve as a link between the GUI library and the rest of the IGSTK toolkit.

Spatial Objects (Geometrical Representation)

Spatial objects define a common structure for geometrical objects defined in IGSTK.  The spatial objects maintain types and physical location of objects in surgical environment.  IGSTK spatial object encapsulate ITK spatial objects in restrictive API.

IGSTK provide several spatial objects. These include as axes, boxes, cone, cylinder, image, ellipsoid, mesh, tube and vascular network.

Spatial Object representation (Visual Representation)

Spatial object representations characterize the graphical representations of the spatial object. The graphical representations dictate how an object should be displayed on the screen. This will include specifying color, opacity and other rendering properties.

Trackers

The tracker component handles communication of tracking tool with tracking device to get position, orientation and other relevant information. The tracking component encapsulates tracking tool, tracking device and all static and dynamic information associated with tracking.

Readers

Readers bring data into the scene generation and representation process.  The most important readers in IGSTK are DICOM image reader classes. These classes are used to read preoperative and intraoperative scans for surgical planning and guidance. Validity check logics are implemented in these classes to avoid incorrect reading and 3D volume generation. Additional readers are available also for loading mesh and calibration data.

In addition to the above main components, IGSTK has infrastructure and service classes. The infrastructure classes include state machines, events, pulse generator, and real time clock generator classes.  Service classes include loggers, registration and calibration classes. Loggers are useful for post analysis of surgical procedures and recovery from a failure.  Integrating loggers into applications streamlines and expedites the application development process and also provides a suitable framework for application verification and validation. This will tremendously enhance the robustness of the system. The other service class in IGSTK is the registration class. This class computes the transformation matrix between the patient and the image coordinate system or between pre-operative and intra-operative images. Furthermore, calibration class is available to measure the most important points of the tool relative to the position of the tracked element on the tool.

Putting it all together

To demonstrate how the components can be integrated to develop an application, we will discuss below image guided needle biopsy application developed using IGSTK.

Needle biopsy is a medical procedure that is performed to analyze the cause of lump or tumor or any abnormal growth in the body.  An interventional radiologist will insert a needle into the tumor and take sample of tissue for laboratory testing. For accurate sampling, the clinician usually uses preoperative images for guidance.  This is a common procedure for lung, breast, and liver and prostate cancer diagnosis.

In this procedure, preoperative CT image is acquired. Using landmark registration technique, preoperative image is registered to the patient coordinate system. During the biopsy procedure, the display is updated with the needle.

The workflow of this application is outlined below.

- Record patient demographic information such as name.

- Load preoperative CT image.  ( Reader component be used )

- Verify that the DICOM image patient name tag matches with the patient information.

- Identify landmark points in the preoperative image using the mouse. A minimum of three landmark points are required for rigid-body transformation computation.

- Initialize the tracking device

- Identify the corresponding landmarks in the physical body using the tracker pointing device.

- Run registration to compute the transformation matrix from patient to preoperative image.

- Identify entry point and target position for needle path planning

- Start tracking.

When tracking is started, an updated position of the probe is displayed overlaid on the preoperative image.

In this application, a Polaris Vicra optical tracker (Northern Digital, Ontario, Canada) is used.  For testing, an abdominal phantom (CIRS Model 57, Norfolk, Virginia) is used. Figure ?? shows  the user interface, which consists of a control panel and four standardized views: axial, sagittal, coronal, and a 3D view. The four view windows show CT images of the abdominal phantom overlaid the path of the biopsy needle. The green cylinder represents the needle which is tracked by the Polaris Vicra Tracker. The four views automatically re-slice and update the images to show the needle tip position as it moves in the anatomy.

INSERT PICTURE OF THE SCREENSHOT OF THE NEEDLE BASED APPLICATION

INSERT PICTURE OF THE EXAMPLE APPLICATION STATE MACHINE

4. Building IGSTK community

Continued success and viability of an open source toolkit depends on a strong user community support.  High quality design and architectural robustness are significant factors that drive the adoption of open source software. However, for a long term success and timely evolutions, contribution from the user community is equally vital. Beyond the first few years of development supported by funding agencies, the lifetime of the toolkit depends on dedicated volunteers who care about the toolkit.  Hence, as part of the toolkit development effort, toolkit creators should treat building and supporting user community as an important task. Without a user community open source software fades away after few years of existence. With this understanding the IGSTK team has been taking key steps to build a strong IGSTK community.

Availability of a complete and sound documentation encourages users to evaluate the toolkit and contribute modifications and bug fixes. Automated documentation generation techniques have simplified this task in IGSTK. Users have access to the latest API documentation from the dashboard. Furthermore, for detailed description of the toolkit and technical explanation of the components, the IGSTK team has written and published a book [INSERT REFERENCE]. The book is available in a PDF format for a free download at the IGSTK website ( www.igstk.org)

Users’ mailing list has recently been launched. Questions posted in this mailing list are promptly answered by the developers of the toolkit. Although recently launched, we have already been getting several questions related to the example applications and functionalities of the some of the components. Furthermore users post feature requests and suggestions on how to improve of the toolkit. These suggestions are seriously studied by the group and documented in the bug tracker and will be considered in the current phase of the project.

Demonstration of applications developed using the toolkit is another key dissemination effort undertaken by the group.  Needle biopsy application developed using IGSTK was demonstrated at SPIE (International society for optical engineering) Medical Imaging conference in 2006 and 2007 at San Diego, CA [INSERT REFERENCE].  Similarly, at  2007 SMIT (Society of Medical Innovation and Technology) conference, a deck of card robot assisted needle placement application was demonstrated [INSERT REFERENCE] (Insert a figure showing deck of cards robot).  This application was developed using IGSTK. These application demos have helped in introducing the toolkit to the medical research community. Related to this effort, tutorial sessions on IGSTK have been offered at these conferences as part of the medical image analysis course using open-source software [INSERT REFERENCE].

Selection of the appropriate license for the toolkit is also essential in building a strong user community. IGSTK is released under BSD-like license. This license allows unrestricted use including in commercial products.  Case in point for building a strong community is the new collaboration that was recently established with the SINTEF Health Research Center at Norway. The center conducts cutting edge research on developing advanced navigation and visualization technologies for image guided surgery. The group is now actively involved in development of new components of the toolkit.

5. Conclusion